## Inspection - Gandhi
Pre-Dispatch Inspection by the Team for the Materials & Eqipment, Hyderabad
Place of Inspection: |  |  |  | IVFC GANDHI, MCH Center 5th Floor, Gandhi Hospital, Secundarabad | [10 empty] | Inspection Dates:
Sl. No | Item Code | Description of the material | Qty as per tender | Qty offered for Inspection | Unit | Quoted price (with all taxes) (Rs.) | Total Amount (Rs) | Make & Model offered in the tender | Make & Model offered | Serial No | Technical Specifications as per TED | Technical Specifications as per Product | Compliance / Deviation if Any | Location / Room Name | Observations of the Inspection Team | Satisfactory (Yes / No) | Remarks
SPLIT AIR CONDITIONERS
1 | IVF149 | Split AC unit 1.5 TR inverter type 3 Star capable of delivering 18000 BTU/hr and above with high wall mounted indoor unit and outdoor condensing unit Hermetically sealed compressor suitable for operation on 230V, 50Hz, 1Phase AC supply capable of performing cooling dehumidifying air circulating and filtering with cooling and condensing units with 3.5 mts of copper piping, insulation kit and 4 mts of 3 core copper flexible chord , built-in stabilizer Voltage range 160- 264 volts and cordless remote control. Makes : Daikin / Thoshiba / O General at Gandhi Hospital | 1 | 2 | Nos | 48000 | 96000 | Makes: Thoshiba / Daikin / O Genral | Daikin |  |  |  |  | UPS ROOM, ANDROLOGY LAB
2 | IVF175 | Split Air Conditioner 1.0 TR: Supply, transportation and installation of split AC unit 1.0 TR inverter type minimum 3 Star capable of delivering 12000 BTU/hr and above with high wall mounted indoor unit and outdoor condensing unit Hermetically sealed compressor suitable for operation on 230V, 50Hz, 1Phase AC supply capable of performing cooling dehumidifying air circulating and filtering with cooling and condensing units with 3.5 mts of copper piping, insulation kit and 4 mts of 3 core copper flexible chord , built-in stabilizer Voltage range 160- 264 volts and cordless remote control. Makes: Daikin / Thoshiba / O General at Gandhi Hospital | 1 | 0 | Nos |  |  | Makes: Thoshiba / Daikin / O Genral | Daikin
3 | IVF173 | Condensing units - VRF/VRV System 16 HP: Supply, Installation, Nitrogen testing , Vacuuming and Commissioning of air cooled variable refrigerant flow modular type condensing units, each comprising of multiple scroll compressors all inverter driven, full charge of refrigerant gas R-410a including additional gas filling , lubricating oil and all accessories as per the specifications. The condensing units shall be suitable to work on cooling mode. The minimum Energy Efficiency Ratio shall be as per ASHRAE STANDARDS 90.1 -2001 table 6.2.1B . The COP of the system shall be greater than 4 under ARI conditions for 100percentage load . The condensing units shall be suitable for operation on 415 plus or minus 10percentage volts, 50Hz, 3 phase AC power supply and , outdoor units to be mounted on MS platform with suitable Vibration Isolation pads to avoid vibrations. All the supports and angle base frame structure to be duly black painted with synthetic epoxy paint. the condensing units shall be of following capacities: The VRF ODU should be capable of working with total capacity of connected indoor units ranging between 100percentage to 140percentage of ODU capacity . COP minimum 3.5. Make : Daikin / LG / O General / Toshiba. 16 HP At Gandhi Hospital | 1 | 1 | set | 325000 | 325000 | Daikin / LG / O General / Toshiba. | Daikin |  |  |  |  | TERRACE FLOOR
4 | IVF174 | VRF Indoor Units : Hi wall Splits 0.8 TR: Supply, Installation, Testing and Commissioning of variable refrigerant flow modular type indoor units wired/wireless remote suitable for R410a refrigerant comprising of all accessories as per the specifications. The indoor units shall be suitable to work on cooling.The minimum Energy Efficiency Ratio shall be as per ASHRAE STANDARDS 90.1 -2001 table 6.2.1B. The indoor units shall be suitable for operation on 220 plus or minus 6percentage volts, 50Hz, 1 phase AC power supply. 0.8 TR At Gandhi Hospital | 2 | 1 | No | 30700 | 30700 |  | Daikin |  |  |  |  | CRYO ROOM
5 | IVF175 | VRF Indoor Units : Hi wall Splits 1.0 TR: Supply, Installation, Testing and Commissioning of variable refrigerant flow modular type indoor units wired/wireless remote suitable for R410a refrigerant comprising of all accessories as per the specifications. The indoor units shall be suitable to work on cooling.The minimum Energy Efficiency Ratio shall be as per ASHRAE STANDARDS 90.1 -2001 table 6.2.1B. The indoor units shall be suitable for operation on 220 plus or minus 6percentage volts, 50Hz, 1 phase AC power supply. 1.0 TR At Gandhi Hospital | 4 | 4 | No | 30700 | 122800 |  | Daikin |  |  |  |  | COUNSELLING ROOM-1, COUNSELLING ROOM-2, BLOOD COLLECTION & INJECTION ROOM, SEMEN COLLECTION ROOM
6 | IVF176 | VRF Indoor Units 1.5 TR : Hi wall Splits : Supply, Installation, Testing and Commissioning of variable refrigerant flow modular type indoor units wired/wireless remote suitable for R410a refrigerant comprising of all accessories as per the specifications. The indoor units shall be suitable to work on cooling.The minimum Energy Efficiency Ratio shall be as per ASHRAE STANDARDS 90.1 -2001 table 6.2.1B. The indoor units shall be suitable for operation on 220 plus or minus 6percentage volts, 50Hz, 1 phase AC power supply. 1.5 TR At Gandhi Hospital | 4 | 5 | No | 35000 | 175000 |  | Daikin |  |  |  |  | ANDROLOGY LAB, MINOR OT, ULTRA SOUND ROOM, UPS ROOM
7 | IVF177 | Cassette Units 3.0 TR : Supply, Installation, Testing and Commissioning of variable refrigerant flow modular type indoor units wired/wireless remote suitable for R410a refrigerant comprising of all accessories as per the specifications. The indoor units shall be suitable to work on cooling mode.The minimum Energy Efficiency Ratio shall be as per ASHRAE STANDARDS 90.1 -2001 table 6.2.1B. The indoor units shall be suitable for operation on 220 plus or minus 6percentage volts, 50Hz, 1 phase AC power supply. 3.0 TR – 4 way At Gandhi Hospital | 1 | 1 | No | 50000 | 50000 |  | Daikin |  |  |  |  | PRE/POST OPERATIVE WARD
8 | IVF178 | Supply, Installation, Testing and Commissioning of Hard copper REFRIGERANT PIPING of 16swg various sizes with necessary supports, fittings, interconnecting the liquid and suction lines between the aircooled outdoor units and the indoor units. The piping shall be carried out of Hard drawn copper pipes of type K with soldered / brazed socket fittings. Isolating valves shall be provided to isolate each compressor, strainers, sight glass and any other component requiring proper operation and maintenance. Both the suction and liquid lines shall be insulated with 19mm thick closed cell insulation for above 25mm dia , 13mm thick for pipes below 25mm dia . The insulation finished with treated woven glass fiber reinforced covering for inside Piping and including making required holes in the wall and re finishing and water tight of holes after laying holes as per the specifications and approved drawings. At Gandhi Hospital | 1 | 1 | Lot | 400000 | 400000
9 | IVF179 | Supply, Installation , Testing and Commissioning of Refnet joints to be provided in refrigerant pipe line. Material of construction for fittings shall be similar to refrigerant piping. At Gandhi Hospital | 10 | 10 | Set | 6191 | 61910
UPS
1 | IVF036 | 20 KVA UPS: Supply and installation of 20KVA 3ph in and 3ph out on line UPS system IGBT based rectifier with input power factor correction 32 bit microprocessor based with isolation transformers double conversion, SNMP controls, LCD display, Pure sine wave output, with auto static switch three phase input / Three phase output and modular structure for better thermal technology / cooling facility, with warranty up to three years.Makes: APC / Numeric / Emerson / Fuji Electric At Gandhi Hospital | 1 | 1 | Nos | 305000 | 305000 | APC / Numeric / Emerson / Fuji Electric | APC |  | 1. SMF battery backup with 1 hours battery backup with SNMP and isolation transformer, battery rack inter connection links (42AH*20 nos) 2. Rectifier & Inverter with IGBT and DSP Technology. 3. Voltage 400 ± 15% V, 3 Phase, 4 Wire. 4. Frequency : 45 to 55 Hz. 5. Display : Control Panel with LCD Screen 6. Battery : Sealed Maintenance Free Rechargeable. 7. Back up : 1 Hr (60 Mts) 8. Protection : Overload & Short Circuit. 9. Ventilation: Forced Air Cooling. |  |  | UPS ROOM
2 | IVF124 | Batteries : Supply and fixing of 12V 65 AH SMF battery including wire leads. Makes: Quanta / Racket / Exide. At Gandhi Hospital | 20 | 32 | Nos | 20584 | 658688 | Makes: Quanta / Racket / Exide. | Exide |  |  |  |  | UPS ROOM
Lab Equipment
1 | IVF001 | Ovum Aspiration Pumps at Gandhi Hospital | 2 | 2 | Nos | 240000 | 480000 |  |  | 2023050005, 2023050006 | 1.Precision built, regulated, vacuum pump for ovum aspiration 2.Rapid suction response at the needle tip and is able to hold constant vacuum Settings accurately, for long periods 3.Dual Display: Actual pump pressure and set pump pressure. 4.Boost feature to clear blockages in Aspiration needle. 5.Ultra-quiet, vibration-free operation. 6.Foot Pedal allows hands-free operation.7.Power Supply: universal input 100-240 V AC, 50/60 Hz 8.Vacuum : -10 to -500 mm Hg. 9.Dimensions (W x D x H)mm: 200x250x115. 10. FDA approved and should be compatible to fit with all commercially available oocyte recovery set |  |  | EMBRYOLOGY LAB
1.Precision built, regulated, vacuum pump for ovum aspiration 2.Rapid suction response at the needle tip and is able to hold constant vacuum Settings accurately, for long periods 3.Dual Display: Actual pump pressure and set pump pressure. 4.Boost feature to clear blockages in Aspiration needle. 5.Ultra-quiet, vibration-free operation. 6.Foot Pedal allows hands-free operation.7.Power Supply: universal input 100-240 V AC, 50/60 Hz 8.Vacuum : -10 to -500 mm Hg. 9.Dimensions (W x D x H)mm: 200x250x115. 10. FDA approved and should be compatible to fit with all commercially available oocyte recovery set
2 | IVF002 | IVF Test tube Warmers at Gandhi Hospital | 2 | 2 | Nos | 49500 | 99000 |  | JD Medicals | 202360, 202359 |  |  |  | EMBROLOGY LAB & ANDROLOGY LAB
|  | 1. IVF test tube warmer ideal for heating of samples in test tubes requiring a stable heating environment 2. Should have at least 12 chambers either for 14ml or 5ml aspiration tubes 3. The highly visible digital LED display indicates the temperature level in either Celsius or Fahrenheit 4. Temperature may be adjusted and is extremely reliable within the range from ambient to 49.9°C. 5. Digital read out of temperature with control panel for settings 6. Temperature accuracy ±0.20C 7. Easy to clean 8. Rectangular warming blocks to hold 10 tubes of 17mm and 12 tubes of 12mm | [8 empty] | 1. IVF test tube warmer ideal for heating of samples in test tubes requiring a stable heating environment | 1. IVF test tube warmer ideal for heating of samples in test tubes requiring a stable heating environment | Yes - Complies
[11 empty] | 2. Should have at least 12 chambers either for 14ml or 5ml aspiration tubes | 2. Should have at least 12 chambers either for 14ml or 5ml aspiration tubes | Yes - Complies
[11 empty] | 3. The highly visible digital LED display indicates the temperature level in either Celsius or Fahrenheit | 3. The highly visible digital LED display indicates the temperature level in either Celsius or Fahrenheit | Yes - Complies
[11 empty] | 4. Temperature may be adjusted and is extremely reliable within the range from ambient to 49.9°C | 4. Temperature may be adjusted and is extremely reliable within the range from ambient to 49.9°C | Yes - Complies
[11 empty] | 5. Digital read out of temperature with control panel for settings | 5. Digital read out of temperature with control panel for settings | Yes - Complies
[11 empty] | 6. Temperature accuracy ±0.20C | 6. Temperature accuracy ±0.20C | Yes - Complies
[11 empty] | 7. Easy to clean | 7. Easy to clean | Yes - Complies
[11 empty] | 8. Rectangular warming blocks to hold 10 tubes of 17mm and 12 tubes of 12mm | 8. Rectangular warming blocks to hold 10 tubes of 17mm and 12 tubes of 12mm
3 | IVF003 | DICOM compatible fully digital, compact, portable, Ultrasound Scan Machine Colour Doppler at Gandhi Hospital | 2 | 2 | Nos | 1365000 | 2730000 |  | Samsung | S2DXM3HW700031D, S2DXM3HW700006H |  |  |  | Ultrasound Room & IVF OT
|  | 1. DICOM compatible fully digital, compact portable Colour Doppler Ultrasound Scan machine. 1.1. System should be operated for general purpose ultrasound scanner and colour Doppler for Pediatric, Abdomen, OB/GYN, Vascular, Cardiac and small parts, Musculoskeletal uses. 1.2. Unit should be of USFDA and European CE approved. 1.3. Latest generation Electronic Phased array Colour Doppler system with Minimum 1,00,000 or more Electronic independent channels. System should be DICOM ready and capable of being interfaced with HIS/RIS/PACS. 1.4. Should be field upgradable to next generation system on site. All new software should be upgraded free of cost during warranty and CMC period. 1.5. Specifications • 256 grey shades for sharp contrast resolutions • Harmonic Imaging- System should have following modes in harmonic with separate setting for: • Tissue Harmonic or equivalent • Harmonic imaging / Equivalent Technology capability in Convex, Linear, Endocavity and Adult Cardiac. 1.6. Gain control in 2D/M Mode/Colour/Doppler with auto adjustment for additional level of flexibility to image quality control. 1.7. Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes 1.8. Frame rate should be 600 or more 1.9. Frequency processing facility for the transducer is 1.0-13.0 MHz +/- 1 MHz. 1.10. Speckle reduction imaging should be available and it should be user configurable setting to obtain better image without reduction in frame rate. 1.11. Steerable PW/CW in all Phased Array probes. 1.12. High-definition acoustic zoom for enlarging sections of 2D and Colour flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate. 1.13. Modes –2D, M-Mode, Steerable PW/CW Doppler, Colour Doppler, and High Definition Colour flow. 1.14. Monitor should be atleast 17” high-resolution colour Monitor. 1.15. Tilt and Swivel monitor should be able to view in all angles and all light conditions. 1.16. Colour Flow Imaging for 1.17. Increased lateral & spatial resolution. 1.18. Selection of different types of colour maps are required 1.19. Colour flow with capability of automatically picking up colour flow as a function of focal depth 1.20. Cine loop memory should be available, specify the memory available 1.21. High Frame rate review for better clarity of playback images study in slow motion. 1.22. Quad loop with memory for pre and post image comparison. 1.23. Simultaneous display of both live B-Mode and 2D with Color should be possible. 1.24. Speckle reduction imaging user configurable should be possible 1.25. Multiline compounding imaging on convex and linear imaging 1.26. Trapezoidal or panoramic imaging should be possible 1.27. User defined system and application presets for multi-user department. 1.28. Inbuilt/ External CD/DVD and PEN drive for image storage, archiving and retrieval for offline complete analysis. . 1.29. System should have facility for cardiac studies like Tissue Doppler imaging, B Mode , M mode ,Cardiac Measurements for Echo Cardiography. 1.30. Transducer ports - Three or more Application software for abdomen, OBG, Vascular, Cardiac Adult and Pediatric applications. All application packages should be built into the system. 1.31. System should be field upgradable to Real time 3D/4D applications for abdomen and fetal anatomical study. 1.32. Digital Storage and Retrieval – 01 1.33. Colour Doppler System with all application packages and probes should be supplied as mentioned below as standard. a. Broadband Convex probe 2.0-6.0 MHz +/- 1 MHz for abdomen and OBG applications - 01 No. b. Broadband Endocavity probe 3.0-9.0 MHz +/- 1 MHz with more than 120 Degree FOV or more- 02 Nos. 1.34. Accessories as standard : 1. Sony Thermal Printer – 01 2. Latest desktop computer with at least I5 configuration. 1.35. GENERAL INSTRUCITONS TO VENDORS: • All information in the tender document must be supported in the product data sheet. • Compliance statement sheet must quote page number/s as it appears in the product data sheet enclosed the vendor. • Supplier should be able to demonstrate its quoted model when and where required along with quality control programme for system performance. • Supplier must attach the list of installation in leading Diagnostics centres hospitals or in institutions inside India (at least 100). • System should be ISO, CE & FDA certified. | [8 empty] | 1.1. System should be operated for general purpose ultrasound scanner and colour Doppler for Pediatric, Abdomen, OB/GYN, Vascular, Cardiac and small parts, Musculoskeletal uses. | 1.1. System should be operated for general purpose ultrasound scanner and colour Doppler for Pediatric, Abdomen, OB/GYN, Vascular, Cardiac and small parts, Musculoskeletal uses.
[11 empty] | 1.2. Unit should be of USFDA and European CE approved. | 1.2. Unit should be of USFDA and European CE approved.
[11 empty] | 1.3. Latest generation Electronic Phased array Colour Doppler system with Minimum 1,00,000 or more Electronic independent channels. System should be DICOM ready and capable of being interfaced with HIS/RIS/PACS. | 1.3. Latest generation Electronic Phased array Colour Doppler system with Minimum 1,00,000 or more Electronic independent channels. System should be DICOM ready and capable of being interfaced with HIS/RIS/PACS.
[11 empty] | 1.4. Should be field upgradable to next generation system on site. All new software should be upgraded free of cost during warranty and CMC period. | 1.4. Should be field upgradable to next generation system on site. All new software should be upgraded free of cost during warranty and CMC period.
[11 empty] | 1.5. Specifications | 1.5. Specifications
[11 empty] | · 256 grey shades for sharp contrast resolutions | · 256 grey shades for sharp contrast resolutions
[11 empty] | · Harmonic Imaging- System should have following modes in harmonic with separate setting for: | · Harmonic Imaging- System should have following modes in harmonic with separate setting for: | Yes - Complies
[11 empty] | · Tissue Harmonic or equivalent | · Tissue Harmonic or equivalent
[11 empty] | · Harmonic imaging / Equivalent Technology capability in Convex, Linear, Endocavity and Adult Cardiac. | · Harmonic imaging / Equivalent Technology capability in Convex, Linear, Endocavity and Adult Cardiac.
[11 empty] | 1.6. Gain control in 2D/M Mode/Colour/Doppler with auto adjustment for additional level of flexibility to image quality control. | 1.6. Gain control in 2D/M Mode/Colour/Doppler with auto adjustment for additional level of flexibility to image quality control.
[11 empty] | 1.7. Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes | 1.7. Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes
[11 empty] | 1.8. Frame rate should be 600 or more | 1.8. Frame rate should be 600 or more
[11 empty] | 1.9. Frequency processing facility for the transducer is 1.0-13.0 MHz +/- 1 MHz. | 1.9. Frequency processing facility for the transducer is 1.0-13.0 MHz +/- 1 MHz.
[11 empty] | 1.10. Speckle reduction imaging should be available and it should be user configurable setting to obtain better image without reduction in frame rate. | 1.10. Speckle reduction imaging should be available and it should be user configurable setting to obtain better image without reduction in frame rate.
[11 empty] | 1.11. Steerable PW/CW in all Phased Array probes. | 1.11. Steerable PW/CW in all Phased Array probes.
[11 empty] | 1.12. High-definition acoustic zoom for enlarging sections of 2D and Colour flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate. | 1.12. High-definition acoustic zoom for enlarging sections of 2D and Colour flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate.
[11 empty] | 1.13. Modes –2D, M-Mode, Steerable PW/CW Doppler, Colour Doppler, and High Definition Colour flow. | 1.13. Modes –2D, M-Mode, Steerable PW/CW Doppler, Colour Doppler, and High Definition Colour flow.
[11 empty] | 1.14. Monitor should be atleast 17” high-resolution colour Monitor. | 1.14. Monitor should be atleast 17” high-resolution colour Monitor.
[11 empty] | 1.15. Tilt and Swivel monitor should be able to view in all angles and all light conditions. | 1.15. Tilt and Swivel monitor should be able to view in all angles and all light conditions.
[11 empty] | 1.16. Colour Flow Imaging for | 1.16. Colour Flow Imaging for
[11 empty] | 1.17. Increased lateral & spatial resolution. | 1.17. Increased lateral & spatial resolution.
[11 empty] | 1.18. Selection of different types of colour maps are required | 1.18. Selection of different types of colour maps are required
[11 empty] | 1.19. Colour flow with capability of automatically picking up colour flow as a function of focal depth | 1.19. Colour flow with capability of automatically picking up colour flow as a function of focal depth
[11 empty] | 1.20. Cine loop memory should be available, specify the memory available | 1.20. Cine loop memory should be available, specify the memory available
[11 empty] | 1.21. High Frame rate review for better clarity of playback images study in slow motion. | 1.21. High Frame rate review for better clarity of playback images study in slow motion.
[11 empty] | 1.22. Quad loop with memory for pre and post image comparison. | 1.22. Quad loop with memory for pre and post image comparison.
[11 empty] | 1.23. Simultaneous display of both live B-Mode and 2D with Color should be possible. | 1.23. Simultaneous display of both live B-Mode and 2D with Color should be possible.
[11 empty] | 1.24. Speckle reduction imaging user configurable should be possible | 1.24. Speckle reduction imaging user configurable should be possible
[11 empty] | 1.25. Multiline compounding imaging on convex and linear imaging | 1.25. Multiline compounding imaging on convex and linear imaging
[11 empty] | 1.26. Trapezoidal or panoramic imaging should be possible | 1.26. Trapezoidal or panoramic imaging should be possible
[11 empty] | 1.27. User defined system and application presets for multi-user department. | 1.27. User defined system and application presets for multi-user department.
[11 empty] | 1.28. Inbuilt/ External CD/DVD and PEN drive for image storage, archiving and retrieval for offline complete analysis. . | 1.28. Inbuilt/ External CD/DVD and PEN drive for image storage, archiving and retrieval for offline complete analysis. .
[11 empty] | 1.29. System should have facility for cardiac studies like Tissue Doppler imaging, B Mode , M mode ,Cardiac Measurements for Echo Cardiography. | 1.29. System should have facility for cardiac studies like Tissue Doppler imaging, B Mode , M mode ,Cardiac Measurements for Echo Cardiography.
[11 empty] | 1.30. Transducer ports - Three or more Application software for abdomen, OBG, Vascular, Cardiac Adult and Pediatric applications. All application packages should be built into the system. | 1.30. Transducer ports - Three or more Application software for abdomen, OBG, Vascular, Cardiac Adult and Pediatric applications. All application packages should be built into the system.
[11 empty] | 1.31. System should be field upgradable to Real time 3D/4D applications for abdomen and fetal anatomical study. | 1.31. System should be field upgradable to Real time 3D/4D applications for abdomen and fetal anatomical study.
[11 empty] | 1.32. Digital Storage and Retrieval – 01 | 1.32. Digital Storage and Retrieval – 01
[11 empty] | 1.33. Colour Doppler System with all application packages and probes should be supplied as mentioned below as standard. | 1.33. Colour Doppler System with all application packages and probes should be supplied as mentioned below as standard.
[11 empty] | a. Broadband Convex probe 2.0-6.0 MHz +/- 1 MHz for abdomen and OBG applications - 01 No. | a. Broadband Convex probe 2.0-6.0 MHz +/- 1 MHz for abdomen and OBG applications - 01 No.
[11 empty] | b. Broadband Endocavity probe 3.0-9.0 MHz +/- 1 MHz with more than 120 Degree FOV or more- 02 Nos. | b. Broadband Endocavity probe 3.0-9.0 MHz +/- 1 MHz with more than 120 Degree FOV or more- 02 Nos.
[11 empty] | 1.34. Accessories as standard : | 1.34. Accessories as standard :
[11 empty] | 1. Sony Thermal Printer – 01 | 1. Sony Thermal Printer – 01
[11 empty] | 2. Latest desktop computer with at least I5 configuration. | 2. Latest desktop computer with at least I5 configuration.
[11 empty] | 1.35. GENERAL INSTRUCITONS TO VENDORS: | 1.35. GENERAL INSTRUCITONS TO VENDORS:
[11 empty] | · All information in the tender document must be supported in the product data sheet. | · All information in the tender document must be supported in the product data sheet.
[11 empty] | · Compliance statement sheet must quote page number/s as it appears in the product data sheet enclosed the vendor. | · Compliance statement sheet must quote page number/s as it appears in the product data sheet enclosed the vendor.
[11 empty] | · Supplier should be able to demonstrate its quoted model when and where required along with quality control programme for system performance. | · Supplier should be able to demonstrate its quoted model when and where required along with quality control programme for system performance.
[11 empty] | · Supplier must attach the list of installation in leading Diagnostics centres hospitals or in institutions inside India (at least 100). | · Supplier must attach the list of installation in leading Diagnostics centres hospitals or in institutions inside India (at least 100).
[11 empty] | · System should be ISO, CE & FDA certified. | · System should be ISO, CE & FDA certified.
4 | IVF004 | Multipurpose Electro hydraulic with manual override mobile OT Table, with divided leg section suitable for all Gynaecological surgical procedures, complete with 5cm Mattress and corded handset at Gandhi Hospital | 1 | 1 | Nos | 563750 | 563750 |  | DESCO | 202206266557 |  |  |  | MAJOR OT
|  | 2.1. Should be a Electro Mechanical controlled operating table, working range from floor level: 640-1040mm 2.2. Should be adjustable to all essential positions. 2.3. Should be equipped with movement controls at side of the table. Should provide wired remote control handset. 2.4. Should have Frame and bottom made of Stainless Steel 304 material. 2.5. Should have reinforced three section stainless steel top/radiolucent top. 2.6. All movements should be adjustable easily. There should be an internal battery backup in the table with auto charging cut off, with batten- charge level indicator. 2.7. Should have detachable head rest which can be easily adjustable to any desired position, above or below table top. 2.8. Table top can be rotated 360° through base, Trendelenburg: ≥25°-30°, Reversed Trendelenburg: ≥30° 2.9. Head Section tilting up from the Horizontal: ≥20°-30°, Head Section tilting down from the Horizontal: ≥28°-30° 2.10. Back Section Raised from the Horizontal: ≥60°-70 ° 2.11. Leg Section Lowered from the Horizontal: ≥40°-50° 2.12. Kidney Position should be achievable by breaking the table. 2.13. Table-top should be radio-lucent. 2.14. User's interface: Automatic by remote control. 2.15. Physical Characteristics • Dimensions (metric): Table top dimension (1900 mm x 525 mm) ± 15% Table elevation: (640mm -1040 mm) ± 10% • Weight (lbs, kg) : Should be able to bear patient having weight upto 160 kg • Heat dissipation: Should maintain nominal temp and the heat should be disbursed through an cooling mechanism • Energy Source: Protection: Should have over charging cut off, • Accessories, Spare Parts, Consumables a) S. S. Arm Rest 1 No b) Anaesthetic Screen 1 No. c) Lithotomy Leg Holders with Stirr-Ups 1 Set d) Leather Wristlets 1 Set e) Padded Leg Rest (Gutter Type)-2 nos f) Anti-static mattress-2 nos g) Additional Poly-urathrene mattress 50mm thick with velcro attchment. 2.16. Environmental And Departmental Considerations • Atmosphere / Ambiance conditioning, humidity, dust ...) • Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90% • User’s care, Cleaning, Disinfection & Sterility issues 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. 2. Sterilization not required. 2.17. Standards And Safety • Certificates : FDA(US) /CE(EU) and BIS/ CDSCO /ISO 13485:2003; IEC 60601-1 2.18. Pre-installation requirements nature, values, quality, tolerance 2.19. Availability of 5amp socket 2.20. Safety and operation check before handover 2.21. Requirements for sign-off: Certificate of calibration and inspection from the manufacturer 2.22. Additional Points • Remote control for all the movements • Battery backup for 50-60 movements | [8 empty] | 1.1. Should be a Electro Mechanical controlled operating table, working range from floor level: 640-1040mm | 1.1. Should be a Electro Mechanical controlled operating table, working range from floor level: 680-950mm | adequate
[11 empty] | 1.2. Should be adjustable to all essential positions. | 1.2. Adjustable to all essential positions. | Yes - Complies
[11 empty] | 1.3. Should be equipped with movement controls at side of the table. Should provide wired remote control handset. | 1.3. Height adjustable by remote handle, with foot-pedal via system | Yes - Complies
[11 empty] | 1.4. Should have Frame and bottom made of Stainless Steel 304 material. | 1.4. Accessories on both sides clamp on standard stainless steel medical rail | Yes - Complies
[11 empty] | 1.5. Should have reinforced three section stainless steel top/radiolucent top. | 1.5. Radiolucent table top with integrated standard x-ray cassette channels. Withlongitudinal shifting capability create perfect X-Ray access. | Yes - Complies
[11 empty] | 1.6. All movements should be adjustable easily. There should be an internal battery backup in the table with auto charging cut off, with batten- charge level indicator. | 1.6. All movements should be adjustable easily. There should be an internal battery backup in the table with auto charging cut off, with batten- charge level indicator.
[11 empty] | 1.7. Should have detachable head rest which can be easily adjustable to any desired position, above or below table top. | 1.7. Independent adjustable head and leg section | Yes - Complies
[11 empty] | 1.8Table top can be rotated 360° through base, Trendelenburg: ≥25°-30°, Reversed Trendelenburg: ≥30° | 1.8Table top can be rotated 360° through base, Trendelenburg: ≥25°-30°, Reversed Trendelenburg: ≥30°
[11 empty] | 1.9. Head Section tilting up from the Horizontal: ≥20°-30°, Head Section tilting down from the Horizontal: ≥28°-30° | 1.9. Head plate up / dpwn : 30° - 45° | adequate
[11 empty] | 1.10. Back Section Raised from the Horizontal: ≥60°-70 ° | 1.10. Back Section Raised from the Horizontal: ≥60°-70 °
[11 empty] | 1.11. Leg Section Lowered from the Horizontal: ≥40°-50° | 1.11. Leg plate up / down : 5° - 90° | Yes - Complies
[11 empty] | 1.12. Kidney Position should be achievable by breaking the table. | 1.12. Built in kidney-bridge , has no negative effect on X-ray radiolucency and is able to support kidney and gall bladder surgeries efficiently | adequate
[11 empty] | 1.13.Table-top should be radio-lucent. | 1.13.Table-top should be radio-lucent. | Yes - Complies
[11 empty] | 1.14. User's interface: Automatic by remote control. | 1.14. User's interface: Automatic by remote control. | Yes - Complies
[11 empty] | 1.15. Physical Characteristics | 1.15. Physical Characteristics
[11 empty] | · Dimensions (metric): Table top dimension (1900 mm x 525 mm) ± 15% Table elevation: (640mm -1040 mm) ± 10% | · Dimensions (metric): Table top dimension (2000 mm x 500 mm) ± 15% Table elevation: (700mm -950 mm) ± 10% | adequate
[11 empty] | · Weight (lbs, kg) : Should be able to bear patient having weight upto 160 kg | · Weight (lbs, kg) : 250 kg | Yes - Complies
[11 empty] | · Heat dissipation: Should maintain nominal temp and the heat should be disbursed through an cooling mechanism | · Heat dissipation: Should maintain nominal temp and the heat should be disbursed through an cooling mechanism
[11 empty] | · Energy Source: Protection: Should have over charging cut off, | · Energy Source: Protection: Should have over charging cut off,
[11 empty] | · Accessories, Spare Parts, Consumables | · Accessories, Spare Parts, Consumables | Yes - Complies
[11 empty] | a) S. S. Arm Rest 1 No | a) S. S. Arm Rest 1 No | Yes - Complies
[11 empty] | b) Anaesthetic Screen 1 No. | b) Anaesthetic Screen 1 No. | Yes - Complies
[11 empty] | c) Lithotomy Leg Holders with Stirr-Ups 1 Set | c) Lithotomy Leg Holders with Stirr-Ups 1 Set | Yes - Complies
[11 empty] | d) Leather Wristlets 1 Set | d) Leather Wristlets 1 Set
[11 empty] | e) Padded Leg Rest (Gutter Type)-2 nos | e) Padded Leg Rest (Gutter Type)-2 nos
[11 empty] | f) Anti-static mattress-2 nos | f) Anti-static mattress-2 nos | Yes - Complies
[11 empty] | g) Additional Poly-urathrene mattress 50mm thick with velcro attchment. | g) Additional Poly-urathrene mattress 50mm thick with velcro attchment. | Yes - Complies
[11 empty] | 1.16. Environmental And Departmental Considerations | 1.16. Environmental And Departmental Considerations
[11 empty] | · Atmosphere / Ambiance conditioning, humidity, dust ...) | · Atmosphere / Ambiance conditioning, humidity, dust ...)
[11 empty] | · Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances | · Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances
[11 empty] | Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90% | Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%
[11 empty] | · User’s care, Cleaning, Disinfection & Sterility issues | · User’s care, Cleaning, Disinfection & Sterility issues | Yes - Complies
[11 empty] | 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | Yes - Complies
[11 empty] | 2. Sterilization not required. | 2. Sterilization not required.
[11 empty] | 1.17. Standards And Safety | 1.17. Standards And Safety
[11 empty] | · Certificates : FDA(US) /CE(EU) and BIS/ CDSCO /ISO 13485:2003; IEC 60601-1 | · Certificates : FDA(US) /CE(EU) and BIS/ CDSCO /ISO 13485:2003; IEC 60601-2 | Yes - Complies
[11 empty] | 1.18. Pre-installation requirements nature, values, quality, tolerance | 1.18. Pre-installation requirements nature, values, quality, tolerance
[11 empty] | 1.19. Availability of 5amp socket | 1.19. Availability of 5amp socket
[11 empty] | 1.20. Safety and operation check before handover | 1.20. Safety and operation check before handover
[11 empty] | 1.21. Requirements for sign-off: Certificate of calibration and inspection from the manufacturer | 1.21. Requirements for sign-off: Certificate of calibration and inspection from the manufacturer
[11 empty] | 1.22. Additional Points | 1.22. Additional Points
[11 empty] | · Remote control for all the movements | · Remote control for all the movements | Yes - Complies
[11 empty] | · Battery backup for 50-60 movements | · Battery backup for 50-60 movements
5 | IVF005 | LED procedure Lights Double dome at Gandhi Hospital | 1 | 1 | Nos | 312500 | 312500 |  | DESCO | 18343,18351 |  |  |  | MAJOR OT
|  | 3.1. Double dome 3.2. Main dome: 1,60,000 lux , second dome 1,60,000 lux for Uniform intensity and Deeper depth penetration. 3.3. Height adjustments : 600mm 3.4. Action radius : 1850 mm 3.5. Possible movements : Radial , Angular & Axial 3.6. Colour temperature : 4500k and above 3.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. 3.8. Intensity, brightness, contrast and power switch to be made available on handle / wall check 3.9. Focal distance (d1+d2) =0.8 to 1.2m 3.10. Temperature rise on the keep of surgeries to be less than 1 degree 3.11. CR +- approx 95 or more 3.12. 360 degree rotation for both arms 3.13. User’s interface: Manual 3.14. Physical Characteristics 3.15. Heat dissipation: Should maintain nominal temperature and the heat should be disbursed through an cooling mechanism 3.16. Energy Source (electricity, UPS, solar, gas, water, CO2 ....) 3.17. Power Requirements: Recharging unit : Input voltage – 220V-240V AC, 50Hz 3.18. Battery operated: 6-8 hours of inbuilt battery back up. 3.19. Tolerance (to variations, shutdowns) : Voltage +10% , frequency +2% 3.20. Protection: Should have over charging cut off with visual symbol. 3.21. Environmental And Departmental Considerations 3.22. Atmosphere / Ambiance conditioning, humidity, dust ...) 3.23. Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances 3.24. Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90% 3.25. User’s care, Cleaning, Disinfection & Sterility issues 3.26. 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. 3.27. 2. Autoclavable Handle should be provided by bidder 4 Pieces. 3.28. Standards And Safety | [8 empty] | 1.1. Double dome | 1.1. Double dome | Yes - Complies
[11 empty] | 1.2. Main dome: 1,60,000 lux , second dome 1,60,000 lux for Uniform intensity and Deeper depth penetration. | 1.2. Main dome: 1,60,000 lux , second dome 1,60,000 lux for Uniform intensity and Deeper depth penetration. | Yes - Complies
[11 empty] | 1.3. Height adjustments : 600mm | 1.3. Height adjustments : 600mm | Yes - Complies
[11 empty] | 1.4. Action radius : 1850 mm | 1.4. Action radius : 1850 mm | Yes - Complies
[11 empty] | 1.5. Possible movements : Radial , Angular & Axial | 1.5. Possible movements : Radial , Angular & Axial | Yes - Complies
[11 empty] | 1.6. Colour temperature : 4500k and above | 1.6. Colour temperature : 4500k and above | Yes - Complies
[11 empty] | 1.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. | 1.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. | Yes - Complies
[11 empty] | 1.8. Intensity, brightness, contrast and power switch to be made available on handle / wall check | 1.8. Intensity, brightness, contrast and power switch to be made available on handle / wall check | Yes - Complies
[11 empty] | 1.9. Focal distance (d1+d2) =0.8 to 1.2m | 1.9. Focal distance (d1+d2) =0.8 to 1.2m | Yes - Complies
[11 empty] | 1.10. Temperature rise on the keep of surgeries to be less than 1 degree | 1.10. Temperature rise on the keep of surgeries to be less than 1 degree | Yes - Complies
[11 empty] | 1.11. CR +- approx 95 or more | 1.11. CR +- approx 95 or more | Yes - Complies
[11 empty] | 1.12. 360 degree rotation for both arms | 1.12. 360 degree rotation for both arms | Yes - Complies
[11 empty] | 1.13. User’s interface: Manual | 1.13. User’s interface: Manual | Yes - Complies
[11 empty] | 1.14. Physical Characteristics | 1.14. Physical Characteristics | Yes - Complies
[11 empty] | 1.15. Heat dissipation: Should maintain nominal temperature and the heat should be disbursed through an cooling mechanism | 1.15. Heat dissipation: Should maintain nominal temperature and the heat should be disbursed through an cooling mechanism | Yes - Complies
[11 empty] | 1.16. Energy Source (electricity, UPS, solar, gas, water, CO2 ....) | 1.16. Energy Source (electricity, UPS, solar, gas, water, CO2 ....) | Yes - Complies
[11 empty] | 1.17. Power Requirements: Recharging unit : Input voltage – 220V-240V AC, 50Hz | 1.17. Power Requirements: Recharging unit : Input voltage – 220V-240V AC, 50Hz | Yes - Complies
[11 empty] | 1.18.Battery operated: 6-8 hours of inbuilt battery back up. | 1.18.Battery operated: 6-8 hours of inbuilt battery back up. | Yes - Complies
[11 empty] | 1.19. Tolerance (to variations, shutdowns) : Voltage +10% , frequency +2% | 1.19. Tolerance (to variations, shutdowns) : Voltage +10% , frequency +2% | Yes - Complies
[11 empty] | 1.20. Protection: Should have over charging cut off with visual symbol. | 1.20. Protection: It have over charging cut off with visual symbol. | Yes - Complies
[11 empty] | 1.21. Environmental And Departmental Considerations | 1.21. Environmental And Departmental Considerations | Yes - Complies
[11 empty] | 1.22. Atmosphere / Ambiance conditioning, humidity, dust ...) | 1.22. Atmosphere / Ambiance conditioning, humidity, dust ...) | Yes - Complies
[11 empty] | 1.23. Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances | 1.23. Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances | Yes - Complies
[11 empty] | 1.24. Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90% | 1.24. Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90% | Yes - Complies
[11 empty] | 1.25. User’s care, Cleaning, Disinfection & Sterility issues | 1.25. User’s care, Cleaning, Disinfection & Sterility issues | Yes - Complies
[11 empty] | 1.26. 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | 1.26. 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | Yes - Complies
[11 empty] | 1.27. 2. Autoclavable Handle should be provided by bidder 4 Pieces. | 1.27. 2. Autoclavable Handle should be provided by bidder 4 Pieces. | Yes - Complies
[11 empty] | 1.28. Standards And Safety | 1.28. Standards And Safety | Yes - Complies
6 | IVF006 | LED procedure Lights Single dome at Gandhi Hospital | 1 | 1 | Nos | 187500 | 187500 |  | DESCO | 18332 |  |  |  | IVF OT
|  | 6.1. Single Dome, Major 6.2.Intensity control : continuous (1,60,000 Lux) 6.3.Height Adjustment : 600mm 6.4.Action radius : 1850mm 6.5.Possible movements : Radial, Angular and Axial 6.6.Colour temperature : 4500 and above 6.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. Intensity, brightness, contrast and power switch to be made available on handle /wall check 6.8.Temperature rise on the keep of surgeries to be less than 1 degree CR + approx 95 or more 6.9.User’s interface: Manual Page 7 of 18 6.10. Physical Characteristics  Heat Dissipation: Should maintain nominal Temp and the heat should be disbursed through an cooling mechanism  Not portable  Energy Source (electricity, UPS, solar, gas, water, CO2 ....)  Power Requirements : Recharging unit: Input voltage- 220V-240V AC, 50Hz  Battery operated : yes  Protection: Should have over charging cut off with visual symbol  Environmental And Departmental Considerations Atmosphere / Ambiance conditioning, humidity, dust ...)  Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances  Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%  User’s care, Cleaning, Disinfection & Sterility issues 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. 2. Autoclavable Handle should be provided by bidder 4 Pieces. Standards And Safety  Certifications: FDA(US) /CE(EU) and BIS/ CDSCO / ISO 13485:2003; IEC 60601-1,  Training And Installation Pre-installation requirements nature, values, quality, tolerance 1. Availability of 5amp socket 2. Safety and operation check before handover Requirements for sign-off: Certificate of calibration and inspection from the manufacturer  Training of staff (medical, paramedical, technicians) OPTIONAL (Depending upon scope of work order) 1. Training of users on operation and basic maintenance 2. Advanced maintenance tasks required shall be documented. | [8 empty] | 6.1. Single Dome, Major | 6.1. Single Dome, Major | Yes - Complies
[11 empty] | 6.2.Intensity control : continuous (1,60,000 Lux) | 6.2.Intensity control : continuous (1,60,000 Lux) | Yes - Complies
[11 empty] | 6.3.Height Adjustment : 600mm | 6.3.Height Adjustment : 600mm | Yes - Complies
[11 empty] | 6.4.Action radius : 1850mm | 6.4.Action radius : 1850mm | Yes - Complies
[11 empty] | 6.5.Possible movements : Radial, Angular and Axial | 6.5.Possible movements : Radial, Angular and Axial | Yes - Complies
[11 empty] | 6.6.Colour temperature : 4500 and above | 6.6.Colour temperature : 4500 and above | Yes - Complies
[11 empty] | 6.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. Intensity, brightness, contrast and power switch to be made available on handle /wall check | 6.7. LED technology minimum 40,000 hours lamp life. The control panel must have a built in digital hour counter to show the number of hours the light has worked. Intensity, brightness, contrast and power switch to be made available on handle /wall check | Yes - Complies
[11 empty] | 6.8.Temperature rise on the keep of surgeries to be less than 1 degree CR + approx 95 or more | 6.8.Temperature rise on the keep of surgeries to be less than 1 degree CR + approx 95 or more | Yes - Complies
[11 empty] | 6.9.User’s interface: Manual Page 7 of 18 | 6.9.User’s interface: Manual Page 7 of 19 | Yes - Complies
[11 empty] | 6.10. Physical Characteristics  Heat Dissipation: Should maintain nominal Temp and the heat should be disbursed through an cooling mechanism  Not portable  Energy Source (electricity, UPS, solar, gas, water, CO2 ....)  Power Requirements : Recharging unit: Input voltage- 220V-240V AC, 50Hz  Battery operated : yes  Protection: Should have over charging cut off with visual symbol  Environmental And Departmental Considerations Atmosphere / Ambiance conditioning, humidity, dust ...)  Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances  Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%  User’s care, Cleaning, Disinfection & Sterility issues | 6.10. Physical Characteristics  Heat Dissipation: Should maintain nominal Temp and the heat should be disbursed through an cooling mechanism  Not portable  Energy Source (electricity, UPS, solar, gas, water, CO2 ....)  Power Requirements : Recharging unit: Input voltage- 220V-240V AC, 50Hz  Battery operated : yes  Protection: Should have over charging cut off with visual symbol  Environmental And Departmental Considerations Atmosphere / Ambiance conditioning, humidity, dust ...)  Operating condition : Capable of operating continuously in ambient temperature of 10 to 40 deg c and relative humidity of 15 to 90% in ideal circumstances  Storage condition : capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%  User’s care, Cleaning, Disinfection & Sterility issues | Yes - Complies
[11 empty] | 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | 1.Parts of the device that are designed to come into contact with the patient or the operator should either be capable of easy disinfection or be protected by a single use / disposable cover. | Yes - Complies
[11 empty] | 2. Autoclavable Handle should be provided by bidder 4 Pieces. Standards And Safety  Certifications: FDA(US) /CE(EU) and BIS/ CDSCO / ISO 13485:2003; IEC 60601-1,  Training And Installation Pre-installation requirements nature, values, quality, tolerance 1. Availability of 5amp socket 2. Safety and operation check before handover Requirements for sign-off: Certificate of calibration and inspection from the manufacturer  Training of staff (medical, paramedical, technicians) OPTIONAL (Depending upon scope of work order) 1. Training of users on operation and basic maintenance 2. Advanced maintenance tasks required shall be documented. | 2. Autoclavable Handle should be provided by bidder 4 Pieces. Standards And Safety  Certifications: FDA(US) /CE(EU) and BIS/ CDSCO / ISO 13485:2003; IEC 60601-1,  Training And Installation Pre-installation requirements nature, values, quality, tolerance 1. Availability of 5amp socket 2. Safety and operation check before handover Requirements for sign-off: Certificate of calibration and inspection from the manufacturer  Training of staff (medical, paramedical, technicians) OPTIONAL (Depending upon scope of work order) 1. Training of users on operation and basic maintenance 2. Advanced maintenance tasks required shall be documented. | Yes - Complies
7 | IVF007 | IVF Workstation with LAF at Gandhi Hospital | 1 | 1 | Nos | 687500 | 687500 |  | JD Medicals | 202225 |  |  |  | EMBRYOLOGY LAB
|  | 1. 6 x 2 Feet IVF Workstation where both user and product are protected against particle and microbial contamination. IVF workstation should have provision for 2 stereo microscopes, one on left and 1 on right.2. Camera Set, which includes Camera, TV tuner Card, Cables and Connect providing Stereozoom Microscope due to Compatibility.3. Surgical Stainless Steel Heated work surface between 2mm-3m.4. Electrical adjusted front glass window 5. 80-100mm heated glass stage to minimize temperature fluctuations in the media with temperature accuracy ± 0.1°C 6. Light source preferred is LED 7. Noise level between 48-52dBA .Programmable auto-start and shut-down functions of all built in features, including air flow 9. The IVF workstation should have main filters which are high-efficiency HEPA filter class (H14). The filter should system consists of two HEPA filters 10. Built in 21 inch minimal Medical grade LCD Monitor 11. Microscope Fittings with Heating Stage Warmer with Microprocessor based Temperature controller 12. Warming blocks should be part of package 13. Should be constructed in high density steel 14. Should be USFDA/European CE certified - 220V/50-60Hz | [8 empty] | 1. 6 x 2 Feet IVF Workstation where both user and product are protected against particle and microbial contamination. IVF workstation should have provision for 2 stereo microscopes, one on left and 1 on right | 1. 6 x 2 Feet IVF Workstation where both user and product are protected against particle and microbial contamination. IVF workstation should have provision for 2 stereo microscopes, one on left and 1 on right | Yes - Complies
[11 empty] | 2. Camera Set, which includes Camera, TV tuner Card, Cables and Connect providing Stereozoom Microscope due to Compatibility. | 2. Camera Set, which includes Camera, TV tuner Card, Cables and Connect providing Stereozoom Microscope due to Compatibility.
[11 empty] | 3. Surgical Stainless Steel Heated work surface between 2mm-3m | 3. Surgical Stainless Steel Heated work surface between 2mm-3m
[11 empty] | 4. Electrical adjusted front glass window | 4. Electrical adjusted front glass window
[11 empty] | 5. 80-100mm heated glass stage to minimize temperature fluctuations in the media with temperature accuracy ± 0.1°C | 5. 80-100mm heated glass stage to minimize temperature fluctuations in the media with temperature accuracy ± 0.1°C | Yes - Complies
[11 empty] | 6. Light source preferred is LED | 6. Light source preferred is LED
[11 empty] | 7. Noise level between 48-52dBA | 7. Noise level between 48-52dBA | Yes - Complies
[11 empty] | 8. Programmable auto-start and shut-down functions of all built in features, including air flow | 8. Programmable auto-start and shut-down functions of all built in features, including air flow | Yes - Complies
[11 empty] | 9. The IVF workstation should have main filters which are high-efficiency HEPA filter class (H14). The filter should system consists of two HEPA filters | 9. The IVF workstation having main filters which are high-efficiency HEPA filter class (H14). The filter should system consists of two HEPA filters | Yes - Complies
[11 empty] | 10. Built in 21 inch minimal Medical grade LCD Monitor | 10. Built in 21 inch minimal Medical grade LCD Monitor | Yes - Complies
[11 empty] | 11. Microscope Fittings with Heating Stage Warmer with Microprocessor based Temperature controller | 11. Microscope Fittings with Heating Stage Warmer with Microprocessor based Temperature controller
[11 empty] | 12. Warming blocks should be part of package | 12. Warming blocks should be part of package | Yes - Complies
[11 empty] | 13. Should be constructed in high density steel | 13. constructed with high density steel | Yes - Complies
[11 empty] | 14. Should be USFDA/European CE certified - 220V/50-60Hz | 14. Should be USFDA/European CE certified - 220V/50-60Hz
8 | IVF008 | IVF Anti-vibration Table for Procedure | 1 | 1 | Nos | 96250 | 96250 |  | JD Medicals | 202071 |  |  |  | EMBRYOLOGY LAB
|  | 1. Air damped Anti-vibration passive table are designed to meet requirements for all models of inverted Microscopes during ICSI procedure 2. The plate on top should be stainless steel 3. Should be constructed in high density steel with added mass to the table 4. Should be ISO certified 5. Dimensions: 780 x 1200 x 790 mm | [8 empty] | 1. Air damped Anti-vibration passive table are designed to meet requirements for all models of inverted Microscopes during ICSI procedure | 1. Air damped Anti-vibration passive table are designed to meet requirements for all models of inverted Microscopes during ICSI procedure | Yes - Complies
[11 empty] | 2. The plate on top should be stainless steel | 2. The top SS 304 grade and pipe with power coated paint. | Yes - Complies
[11 empty] | 3. Should be constructed in high density steel with added mass to the table | The inner table shall be heavily weighted pre-fabricated | Yes - Complies
[11 empty] | 4. Should be ISO certified | 4. Should be ISO certified
[11 empty] | 5. Dimensions: 780 x 1200 x 790 mm | 5. Dimensions: 780 x 1200 x 790 mm
9 | IVF009 | CO2 Incubators | 2 | 2 | Nos | 825000 | 1650000 |  | ThermoFisher | 43009230, 43008170 |  |  |  | EMBRYOLOGY LAB
|  | 1. 160- 220 liters internal Volume 2. Stainless steel interior 3. TC/IR CO2 Sensor Soft touch keypad 4. Out Door Heating Error Warning 5. Auto Start Function 6. Inner Safety Glass Door 7. HEPA Filter for CO2 8. Inner safety glass screen ( 3-4 pieces) 9. Microprocessor PID temperature controller 10. Temperature Sensor made up of Platinum 11. Temperature range ambient above 5°C to 50°C ( ± 0.1°C at 37°C) 12. Air- Jacket direct heating system 13. Over-Temperature Protection Smooth inner casing with rounded corners on all sides made of stainless steel Dry Or Moist heating disinfect ion at 90°C to 120°C. 14. Disinfection duration 12-13 hrs 15. 3-4 sets of shelves. Perforated Shelves for uniform heat distribution/ shelf adjustment.16. LED/LCD Display of chamber temp. and CO2 level. CO2 Control range: 0 to temperature alarm 32°C-47°C. 17. Humidity: Active sterile humidity control through vaporizing module operating at 120ºC Measuring range 0-98% RH /Range 60-95% RH. The system should have easy-to-set high-low humidity levels 18. The system should have a built-in decontamination cycle for the complete elimination of bacteria, fungi, spores, mycoplasma, etc. The system must have onboard graphic capability /data logging to enable users to obtain historical performance. The incubator must have a fully automatic start routine function.19. Must be FDA/EU approved | [8 empty] | 1. 160- 220 liters internal Volume | 1. 165 liters internal Volume | Yes - Complies
[11 empty] | 2. Stainless steel interior | 2. electropolished Stainless steel interior | Yes - Complies
[11 empty] | 3. TC/IR CO2 Sensor Soft touch keypad | 3. Touch screen interface | adequate
[11 empty] | 4. Out Door Heating Error Warning | 4. Out Door Heating Error Warning
[11 empty] | 5. Auto Start Function | 5. Auto Start Function
[11 empty] | 6. Inner Safety Glass Door | 6. Inner Safety Glass Door
[11 empty] | 7. HEPA Filter for CO2 | 7. HEPA Filter for CO2 | Yes - Complies
[11 empty] | 8. Inner safety glass screen ( 3-4 pieces) | 8. Inner safety glass screen ( 3-4 pieces)
[11 empty] | 9. Microprocessor PID temperature controller | 9. Microprocessor PID temperature controller | Yes - Complies
[11 empty] | 10. Temperature Sensor made up of Platinum | 10. Temperature Sensor made up of Platinum
[11 empty] | 11. Temperature range ambient above 5°C to 50°C ( ± 0.1°C at 37°C) | 11. Temperature range ambient above 3°C to 55°C ( ± 0.1°C at 37°C) | Yes - Complies
[11 empty] | 12. Air- Jacket direct heating system | 12. Air- Jacket direct heating system
[11 empty] | 13. Over-Temperature Protection Smooth inner casing with rounded corners on all sides made of stainless steel Dry Or Moist heating disinfect ion at 90°C to 120°C. | 13. Over-Temperature Protection Smooth inner casing with rounded corners on all sides made of stainless steel Dry Or Moist heating disinfect ion at 90°C to 120°C.
[11 empty] | 14. Disinfection duration 12-13 hrs | 14. Disinfection duration 12-13 hrs
[11 empty] | 15. 3-4 sets of shelves. Perforated Shelves for uniform heat distribution/ shelf adjustment. | 15. 3-4 sets of shelves. Perforated Shelves for uniform heat distribution/ shelf adjustment.
[11 empty] | .16. LED/LCD Display of chamber temp. and CO2 level. CO2 Control range: 0 to temperature alarm 32°C-47°C | .16. LED/LCD Display of chamber temp. and CO2 level. CO2 Control range: 0 to temperature alarm 32°C-47°C | Yes - Complies
[11 empty] | 17. Humidity: Active sterile humidity control through vaporizing module operating at 120ºC Measuring range 0-98% RH /Range 60-95% RH. The system should have easy-to-set high-low humidity levels | 17. Humidity: Active sterile humidity control through vaporizing module operating at 120ºC Measuring range 0-98% RH /Range 60-95% RH. The system should have easy-to-set high-low humidity levels
[11 empty] | 18. The system should have a built-in decontamination cycle for the complete elimination of bacteria, fungi, spores, mycoplasma, etc. The system must have onboard graphic capability /data logging to enable users to obtain historical performance. The incubator must have a fully automatic start routine function | 18. our advanced contamination technologies are designed to protect your valuable cultures, eliminate the loss of time and resources while providing convenient added security for your research work. | adequate
[11 empty] | 19. Must be FDA/EU approved | 19. CE certified | adequate
10 | IVF010 | Trigas Bench Top Incubators for Human Embryo culture | 1 | 1 | Nos | 825000 | 825000 |  | PLANER | 55339 |  |  |  | EMBRYOLOGY LAB
|  | 1. Compact, humidified Pre-Mixed/ TRI-gas incubator designed to maintain optimal environment for development of ova or embryos.2. Has two or more chambers to hold dishes, dishes per chamber minimum 4.3. The dual chambers have individual temperature control. Heating of both base and lid ensures even distribution of heat.4. Direct contact of base with each culture dish ensures heat transfer by Conduction resulting in thermal stability.5. Automatic gas purge on lid closure to maintain gas environment. 6. In built humification system.7. Uses pre-mix Gas & Several units can be connected to one source of Supply.8. 24 hour digital recording of temperature and gas flow.9. Time-stamped alarm notifications include description of event.10. Battery backup - For day 1 - day 5 Culturing 11. Gas Supply: High Purity CO2/O2/N2 mixture. Nominal input pressure 150kPa. 12. Power: Universal Input 100-240 V AC, 50/60 Hz. 13. Safety : Designed to conform with AS3200.1 1990, IEC60601.1 and IEC61010.1 14. Two Stage Regulator 15. Product should be approved from USFDA/European CE certified. | [8 empty] | 1. Compact, humidified Pre-Mixed/ TRI-gas incubator designed to maintain optimal environment for development of ova or embryos. | 1. Compact, humidified Pre-Mixed/ TRI-gas incubator designed to maintain optimal temperature, gas and humidity | Yes - Complies
[11 empty] | 2. Has two or more chambers to hold dishes, dishes per chamber minimum 4 | 2. Large range of cultural dish sizes accepted. | adequate
[11 empty] | 3. The dual chambers have individual temperature control. Heating of both base and lid ensures even distribution of heat. | 3. Independent temperature monitoring | Yes - Complies
[11 empty] | 4. Direct contact of base with each culture dish ensures heat transfer by Conduction resulting in thermal stability | 4. Direct contact of base with each culture dish ensures heat transfer by Conduction resulting in thermal stability
[11 empty] | 5. Automatic gas purge on lid closure to maintain gas environment. | 5. Minimal tubing to optimal temperature, gas and humidity | adequate
[11 empty] | 6. In built humification system. | 6. In built humification system. | Yes - Complies
[11 empty] | 7. Uses pre-mix Gas & Several units can be connected to one source of Supply. | 7. Uses pre-mix Gas & Several units can be connected to one source of Supply.
[11 empty] | 8. 24 hour digital recording of temperature and gas flow. | 8. 24 hour digital recording of temperature and gas flow.
[11 empty] | 9. Time-stamped alarm notifications include description of event. | 9. Time-stamped alarm notifications include description of event.
[11 empty] | 10. Battery backup - For day 1 - day 5 Culturing | 10. Battery backup - For 2 hours | adequate
[11 empty] | 11. Gas Supply: High Purity CO2/O2/N2 mixture. Nominal input pressure 150kPa. | premixed gas. Typically 6% CO2, 5% O2, 89% N2 | adequate
[11 empty] | 12. Power: Universal Input 100-240 V AC, 50/60 Hz. | 12. Power: Universal Input 100-240 V AC, 50/60 Hz. | Yes - Complies
[11 empty] | 13. Safety : Designed to conform with AS3200.1 1990, IEC60601.1 and IEC61010.1 | 13. Safety : Designed to conform with AS3200.1 1990, IEC60601.1 and IEC61010.1
[11 empty] | 4. Two Stage Regulator | 4. Two Stage Regulator
[11 empty] | 15. Product should be approved from USFDA/European CE certified. | 15. Product should be approved from USFDA/European CE certified.
11 | IVF011 | Trinocular Stereozoom Microscope | 1 | 1 | Nos | 309000 | 309000 |  | NIKON |  |  |  |  | EMBRYOLOGY LAB
|  | 1. Microscope body with 5x-500x magnification, and 45 degrees inclination trinocular tube (C- mount) 2. Eyepiece 10X with ESD capability, F.N. 22, focusable 3. Focussing mount with coaxial coarse and fine focussing knobs 4. Power Cord 5. Should be integrated into IVF Workstation Dust Cover`6. Beam sp,litter for trinocular port 7. Eyepiece 10x, FN 22 8. C-mount adapter to be included 9. Trinocular observation tube with inclination at 30 degree, inter pupillary distance adjustment 48mm -76mm. 10. Working Distance Up to 92 mm 11. Plane achromatic objective 1X Resolution of at least 600lp/mm.12. Transmitted light stage with halogen illuminator of at least 30 watt.13. Scientific Digital Camera for microscopy with control software, progressive scan CMOS/CCD having resolution of 5 M-pixel or better14. Coaxial course and fine focusing knob mechanism should be built in | [8 empty] | 1. Microscope body with 5x-500x magnification, and 45 degrees inclination trinocular tube (C- mount) | 1. Microscope body with 5x-500x magnification, and 45 degrees inclination trinocular tube (C- mount)
[11 empty] | 2. Eyepiece 10X with ESD capability, F.N. 22, focusable | 2. Eyepiece 10X with ESD capability, F.N. 22, focusable | Yes - Complies
[11 empty] | 3. Focussing mount with coaxial coarse and fine focussing knobs | 3. Focussing mount with coaxial coarse and fine focussing knobs | Yes - Complies
[11 empty] | 4. Power Cord | 4. Power Cord | Yes - Complies
[11 empty] | 5. Should be integrated into IVF Workstation Dust Cover` | 5. Should be integrated into IVF Workstation Dust Cover`
[11 empty] | 6. Beam sp,litter for trinocular port | 6. Beam sp,litter for trinocular port
[11 empty] | 7. Eyepiece 10x, FN 22 | 7. Eyepiece 10x, FN 23 | Yes - Complies
[11 empty] | 8. C-mount adapter to be included | 8. C-mount adapter to be included
[11 empty] | 9. Trinocular observation tube with inclination at 30 degree, inter pupillary distance adjustment 48mm -76mm. | 9. Trinocular observation tube with inclination at 30 degree, inter pupillary distance adjustment 48mm -76mm.
[11 empty] | 10. Working Distance Up to 92 mm | 10. Working Distance Up to 92 mm
[11 empty] | 11. Plane achromatic objective 1X Resolution of at least 600lp/mm. | 11. Plane achromatic objective 1X Resolution of at least 600lp/mm.
[11 empty] | 12. Transmitted light stage with halogen illuminator of at least 30 watt. | 12.there are two illuminators available: the G-LS of 6V10W and the C-DSLS of 6V20W | adequate
[11 empty] | 13. Scientific Digital Camera for microscopy with control software, progressive scan CMOS/CCD having resolution of 5 M-pixel or better | 13. Scientific Digital Camera for microscopy with control software, progressive scan CMOS/CCD having resolution of 5 M-pixel or better
[11 empty] | 14. Coaxial course and fine focusing knob mechanism should be built in | 14. Coaxial course and fine focusing knob mechanism should be built in | Yes - Complies
12 | IVF012 | ICSI Micro Manipulator with Inverted Microscope | 1 | 1 | Nos | 2625000 | 2625000 |  | NIKON | 2511 - 753037 |  |  |  | EMBRYOLOGY LAB
|  | 1. Latest Inverted microscope 2. 3 position light path 3. Prism preinstalled 4. Illumination equipped with condenser holder 5. LED Lamphouse 6. Sextuple Nosepiece is included.7. Binocular Tube Widefield eyepiece 10X, focusable.8. Enough travelling range applied for slide glass, 35 mm dish as well as multiwell plate, circle stage inserts included.9. Sample holder for slide glass, 35 mm dish, multiwell plate as option. Stage stopper function is implemented for time-lapse or operation on stage.10. Long working distance condenser for DIC/RC/NAMC observation11. Mechanical or manual or electrical micromanipulator12. Digital display of 3D fine control from single level 13. Heating system of insert should be such that it should avoid cold spots 14. 1 x Oil Injector 15. 1 x air injector 16. For Intracyloplasmic Sperm Injection in IVF. 17. Micromanipulator should be compatible for the utility of commonly available LASER And Spindle view imaging system.18. Must be FDA/ EU approved | [8 empty] | 1. Latest Inverted microscope | 1. Latest Inverted microscope | Yes - Complies
[11 empty] | 2. 3 position light path | Motion sensor LED light underneath the stage for improved visibility when changing objectives. | adequate
[11 empty] | 3. Prism preinstalled | 3. Prism preinstalled | Yes - Complies
[11 empty] | 4. Illumination equipped with condenser holder | 4. Illumination equipped with condenser holder
[11 empty] | 5. LED Lamphouse | 5. LED Lamphouse
[11 empty] | 6. Sextuple Nosepiece is included. | 6. Sextuple Nosepiece is included.
[11 empty] | 7. Binocular Tube Widefield eyepiece 10X, focusable. | 7. Binocular Tube Widefield eyepiece 10X, focusable.
[11 empty] | 8. Enough travelling range applied for slide glass, 35 mm dish as well as multiwell plate, circle stage inserts included | 8. Enough travelling range applied for slide glass, 35 mm dish as well as multiwell plate, circle stage inserts included
[11 empty] | 9. Sample holder for slide glass, 35 mm dish, multiwell plate as option. Stage stopper function is implemented for time-lapse or operation on stage | 9. Sample holder for slide glass, 35 mm dish, multiwell plate as option. Stage stopper function is implemented for time-lapse or operation on stage
[11 empty] | .10. Long working distance condenser for DIC/RC/NAMC observation | .10. Long working distance condenser for DIC/RC/NAMC observation
[11 empty] | 11. Mechanical or manual or electrical micromanipulator | 11. Mechanical or manual or electrical micromanipulator
[11 empty] | 12. Digital display of 3D fine control from single level | 12. 3.5" color LCD touch screen display for temperature control. | adequate
[11 empty] | 13. Heating system of insert should be such that it should avoid cold spots | 13. Thermosafe heating system with heated metal insert: 25mm aperture (standard), 16mm aperture, ITO heated glasss insert | Yes - Complies
[11 empty] | 14. 1 x Oil Injector | 14. 1 x Oil Injector
[11 empty] | 15. 1 x air injector | 15. 1 x air injector
[11 empty] | 16. For Intracyloplasmic Sperm Injection in IVF | 16. For Intracyloplasmic Sperm Injection in IVF
[11 empty] | 17. Micromanipulator should be compatible for the utility of commonly available LASER And Spindle view imaging system. | 17. Micromanipulator should be compatible for the utility of commonly available LASER And Spindle view imaging system.
[11 empty] | .18. Must be FDA/ EU approved | .18. Must be FDA/ EU approved
13 | IVF013 | IVF Laser System | 1 | 1 | Nos | 2560000 | 2560000 |  | NIKON | 21086 |  |  |  | EMBRYOLOGY LAB
|  | 13.1. Laser for IVF lab to ablate zona pellucida, hatching, biopsies and blastomere collapse 13.2. Moveable laser with biopsy mode 13.3. The laser objective should be designed in such a way that it can focus visible light,but also to focus the infra-red beam at the same plane as the visible light, and to maximise the power transmission for efficient drilling 13.4. Laser should include both pilot and ablation laser 13.5. Laser should have a option to indicate heat safety zone 13.6. Laser power should be between 350-425mW 13.7. Powerful software should be supplied with laser to ensure that it can have home position setting, measuring tool digital magnification, simulator for training and demo features. | [8 empty] | 13.1. Laser for IVF lab to ablate zona pellucida, hatching, biopsies and blastomere collapse | 13.1. Laser for IVF lab to ablate zona pellucida, hatching, biopsies and blastomere collapse
[11 empty] | 13.2. Moveable laser with biopsy mode | 13.2. Moveable laser with biopsy mode
[11 empty] | 13.3. The laser objective should be designed in such a way that it can focus visible light,but also to focus the infra-red beam at the same plane as the visible light, and to maximise the power transmission for efficient drilling | 13.3. The laser objective should be designed in such a way that it can focus visible light,but also to focus the infra-red beam at the same plane as the visible light, and to maximise the power transmission for efficient drilling
[11 empty] | 13.4. Laser should include both pilot and ablation laser | 13.4. Laser should include both pilot and ablation laser
[11 empty] | 13.6. Laser power should be between 350-425mW | 13.6. Laser power between 1480nm / 400mW | Yes - Complies
[11 empty] | 13.7. Powerful software should be supplied with laser to ensure that it can have home position setting, measuring tool digital magnification, simulator for training and demo features. | 13.7. Powerful software should be supplied with laser to ensure that it can have home position setting, measuring tool digital magnification, simulator for training and demo features. | Yes - Complies
14 | IVF014 | IVF Thermometer with Probe | 1 | 1 | Nos | 61875 | 61875 |  | Testo | 84231125 | 1. Easy to use handheld digital, water proof 2. To check temperature of heated surface, inside the media droplet, incubators, refrigerators 3. Operating time at least 200 hours on battery 4. Ergonomic design and should have digital display 5. Calibration certificate required. 6. Certificate: should be CE or FDA or BIS certified |  |  | EMBROLOGY LAB
15 | IVF015 | CO2 Cylinder with accessories | 2 | 2 | Nos | 50000 | 100000 |  | Everest Kanto | L084,J063 | 1. Medical grade CO2 cylinder 31 Kg. Accessories like pliers, wrenches to work with each cylinders to be provided, 06 extra regulators for each cylinder. 2. With two stage regulator calibrated in PSI 3. With required attachments for manifold. With automatic inline arrangement to ensure steady, nonstop supply |  |  | MANIFOLD ROOM
16 | IVF016 | CO2 and O2 Analyzer | 1 | 1 | Nos | 378000 | 378000 |  | Geotech | IN16815 | 1. Hand held with Quick verification of CO2 and O2 incubator levels 2. Large data storage up to 750 readings 3. CO2 0-20%; O2 0-100% 4. Battery Life: 10 Hours (8 hours pumping) 5. Easy to read 6. Atleast 2 temperature probes 7. Built in moisture gas removal and humidity trap |  |  | EMBRYOLOGY LAB
17 | IVF017 | PH Meter | 1 | 1 | Nos | 93750 | 93750 |  | THERMO SCIENTIFIC | 3071405 | 1. Hand held, simple pH meter to read the media drops 2. pH accuracy ±0.03pH 3. pH calibration atleast 2 points in room temperature 4. Calibration by pH buffers 7-10 5. Temperature accuracy ±0.50C 6. Data logging facility |  |  | EMBRYOLOGY LAB
18 | IVF018 | VOC Meter for ART Lab | 1 | 1 | Nos | 731250 | 731250 |  |  | 592931054 | 1. Handheld monitor capable of detecting contamination at 0.1ppm 2. PID sensor 0-15000ppm 3. Response time < 5 seconds 4. Flow rate 400-500cc/min 5. Data logging facility mandatory 6. Weight less then 1Kg |  |  | EMBRYOLOGY LAB
19 | IVF019 | Sperm Counting Chamber | 1 | 1 | Nos | 75625 | 75625 |  | Shukratara |  | 1. Counting Chamber with cover Slip(Glass) 2. Cover Slip with Grid built in with 100 squares 3. Reusable 4. No Calibration required 5. Optimal depth 10 microns 6. Should be provided with cleaning brush and cleaning paper 7. 99% accuracy for pre and post wash semen analysis 8. Valid CE/US FDA/BIS Certification 9. NO DILUTION REQUIRED FOR SPERM COUNT CALCULATION |  |  | ANDROLOGY LAB
20 | IVF020 | Binocular Microscope MLX-B Plus | 1 | 1 | Nos | 41250 | 41250 |  | Magnus | 23F0212 | 1. Standard microscope set complete with built-in 6V20W halogen light illuminator with regular power supply 2. Quadruple ball bearing nosepiece, co-axial coarse and fine focusing controls, high resolution Semi- Plan Achromat objectives 4x, 10x, 20X & 40x (spring) 3. 3600 rotatable inclined binocular tube, fungus resistant optics for tropical use, widefield paired eyepiece HWF10x (F.N.18), right hand control co-axial low drive mechanical stage and rack & pinion focusable Abbe condenser 0.9/1.25 N.A with iris diaphragm complete set in thermocole packing |  |  | ANDROLOGY LAB
21 | IVF021 | Laminar Flow Cabinet with Andrology Work Station | 1 | 1 | Nos | 250000 | 250000 |  |  | 20130 | 1. 3 Feet LAF where both user and product are protected against particle and microbial contamination. IVF2. Quadruple ball bearing nosepiece, co-axial coarse and fine focusing controls, high resolution Semi- Plan Achromat objectives 4x, 10x, 20X & 40x (spring)3. 3600 rotatable inclined binocular tube, fungus resistant optics for tropical use, widefield paired eyepiece HWF10x (F.N.18), right hand control co-axial low drive mechanical stage and rack & pinion focusable Abbe condenser 0.9/1.25 N.A with iris diaphragm complete set in thermocole packing4. Surgical Stainless Steel (2mm) work surface5. Toughened glass front window6. Adjustable luminosity7. Fully operated with a microprocessor control panel.8. USB data output 9. Electrical outlets10. Fixed support Stand (80-85cm)11. Dimensions: 1000x800x2000 (WxDxH) |  |  | ANDROLOGY LAB
22 | IVF022 | Clinical Centrifuge machine for Andrology with Acessories | 1 | 1 | Nos | 34500 | 34500 |  | REMI | ZJGN20240 | 1. Brushless Motor Capacity (6x10ml) with Swing Out Rotor 2. Max RPM/RCF(xg)-4000RPM/2270g based on the rotor 3. Speed setting as 500-4000 RPM in steps of 10RPM 4. User should be able to set and save upto 99 user defined programs(protocols) with a digital display |  |  | ANDROLOGY LAB
23 | IVF023 | DRY Incubator | 1 | 1 | Nos | 563750 | 563750 |  | ThermoFisher | 30571816 | As per Annexure - I |  |  | ANDROLOGY LAB
24 | IVF024 | Pipetter and Denudation system | 1 | 1 | Nos | 103125 | 103125 |  |  | 148956 | 1. Denudation System It should be a medical device for use in Human IVF.2. Each system should consist of following items • Denudation Pipette Rack Should have slots for placing at least two denudation holder with pipette. • Denudation Pipette Holder • Adjustable handle to accept all sizes of pipettes 3. Denudation Pipette• Flexible polycarbonate pipettes used for manipulation of oocytes and embryo Pipette tips should be suitable to get easily attached & detached with holder Size: 135-140 µm, pack of 10 or 20 (5 Packs)• Size:170-175 µm, pack of 10 or 20 (5 Packs) • Size: 275-300µm, pack of 10 or 20 (5 Packs)4. Certificate : Should be European CE or US FDA or BIS certified. |  |  | EMBRYOLOGY LAB
25 | IVF025 | Pharmaceutical Refrigerator | 1 | 1 | Nos | 225000 | 225000 |  |  | 210330355 | 1. Capacity of storage 300 liters or more2. Temp range-should have adjustable temperature control range from +1° to +8° C, factory present at 4° C.3. Refrigerator system: The system should have high density CFC-free insulation to protect cabinet from ambient temperature fluctuation. The system should have positive, force, air circulation to maintain temperature uniformity at all shelf levels, with quick recovery +/- 1 degree Celsius.4. The system should have sensors for activation of utomatic/manual defrosts cycle to minimize the frost build up.5. Internal construction should be made of high grade stainless steel (minimum 22G) External construction corrosion resistant sheet at least 1 mm thickness.6. Internal temperature control: System should have temperature control range from +1O C to +8O C. Temperature control resolution should be better than 1O C.Cooling down time of max of 150 min on half load.7. External ambient temp should perform in ambient temp up to +43O C.8. Door System should lockable double doors with double pane with self closing door for better safety. |  |  | ANDROLOGY LAB
26 | IVF026 | Cryocane TA 26 Ltr with wheels | 2 | 2 | Nos | 61875 | 123750 |  | IBP | V23300, J-23589 | 1. For specimen storage - 47 Litres-1 pcs 2. For Liquid Nitrogen storage - 26 Litres – 1 pcs |  |  | CRYO ROOM
27 | IVF027 | Defibrillator | 1 | 1 | Nos | 312500 | 312500 |  | Nihon Kohden | 10762 | As per Annexure - I |  |  | MAJOR OT
28 | IVF028 | 7 Para Patient Monitor | 2 | 2 | Nos | 280000 | 560000 |  |  | E7230826021, E7230826024 | 1.Description of function: Beside monitors are used to monitor the seven vital parameters (ECG,SPO2,NIBP,Respiration ,Temperature ,Dual IBP AND Etc02) of patient’s side inwards ,ICU ,CCU , other intensive care units and all operation theaters.2. Operational Requirements: patient monitors must be pre-configured with following seven parameters ECG,SP02,NIBP,Respiration,Temparature ,Dual IBP, and Etc02 ) easy to use, portable ,wall mounted and operation by single knob control.3.Technical specifications: patient monitor should be of modular design so that software and hardware up gradation can be done on the client location a) patient monitor must have Bidirectional connectivity to central monitoring system with other bed view facility. b) Deletedc)protection: built -in cautery & defibrillator protection d) minimum 72 hours tabular trends must be available for all monitors.4. Display : 1) Monitors must have at least 12.1 inch or more high resolution TFT colour display screen with flexible user definable settings2) should have resolution of 800×600 or better with at least 8 traces and numeric values on display 3) combination of single , dual and multi parameters modes 4) should have Ability to change the wave trace color and parameter colour by user 5) Must have special features to adjust and smooth the wave forms for better viewing from from different angle 6) must have RJ 45/RS232 interface for CMS ( central monitoring system) 5. ECG:Monitoring of ECG (I,II,III,aVR,avL,aVL,aVF and chest lead) and pulse detection ,display of heart rate with low and high heart rate alarm ( adjustable between 30-250 bpm, audio visual alarms) |  |  | IVF OT & MAJOR OT
29 | IVF029 | Suction Machine | 1 | 1 | Nos | 49500 | 49500 |  |  | 00168 | 1. Portable suction unit: size 38x17x28.5cm, weight 5.1kg 2. Technical specifications: SUCTION MACHINE with pump: • Power supply: 230-240V/50Hz• Vacuum capacity: 18 litres/mim • Maximum depression: -75kPa (-563mmHg)• Vacuum is created by a plastic piston and cylinder system, with four vacuum-creating modules (QuatroFlex™ technology).• The membrane vacuum regulator permits accurate vacuum settings.• Double overflow-protection system (bottle and pump).• Working temperature range: +5 to +40°C.• SAFETY DEVICE:The safety device protects the pump against overflow: suction is interrupted when foam or liquids reach the safety level in |  |  | MAJOR OT
30 | IVF030 | Anaesthesia Work station at Gandhi Hospital | 1 | 1 | Nos | 937500 | 937500 |  | Comen | X4230427001B | As per Annexure - I |  |  | MAJOR OT
31 | IVF031 | Autoclave at Gandhi Hospital | 1 | 1 | Nos | 562500 | 562500 |  | Cistron | CSHSMS023 | 1. High Grade strong stainless steel, Triple walled construction.2. Positive radial self-locking safety doors.3. Hydrostatically tested to withstand 2.5 times the working pressure.4. Sealed with Neoprene/Silicon long-lasting and durable gasket.5. Digital display for Jacket and Chamber pressure and temperature.6. Outer jacket insulated to prevent heat loss;with a high grade insulationmaterial 7. Mounted on 304 stainless steel frame with ground leveling flanges.8. Temperature and pressure cut-off device.9. Auto cut-off at low water level10. Rust-proof 304 grade stainless steel.11. Cylindrical construction.12. Equipment should have separate steam release valve and drainage system.13. Minimum of two safety valves with auto-release at 16 and 20.14. User’s interface : Manual |  |  | AUTOCLAVE ROOM
32 | IVF032 | Operative Hysteroscopy and Resectoscope Set Machine | 1 | 1 | Nos | 2125000 | 2125000 |  |  | 232409075W | 1. 2.9 mm hysteroscope 2. "Hysteroscope sheath for 2.9 mm telescope — operative inner sheath 3. "Hysteroscope sheath for 2.9 mm telescope — operative outer sheath 4. Accessories for hysteroscope — • Hysteroscopy scissor • Hysteroscopy grasper • Hysteroscopy biopsy forceps • Monopolar electrode • Bipolar electrode • Monopolar cable • Bipolar cable • Hysteroscopy pressure bag - 1000ml • Hysteromat • Vasa Point |  |  | MAJOR OT
33 | IVF033 | Operative Laparoscopy Set | 1 | 1 | Nos | 4200000 | 4200000 |  |  | 232409075W | Telescope 2.9 mm, Telescope 5 mm, Telescope 10 mm, Lap Monopolar Hook, Lap Bipolar Insert, Suction cannula, Hys Inner and outer sheath, Medical Grade Monitor, Camera set with cable, Light Source with cable, insufflator and gas tube, Suture removal scissor, L- Hook and Insert, Lap suction set, Lap mariland, Lap Bowl Grasper, Lap scissor, Lap Allies, Lap needle holder, Lap PCOD drilling, Mayoma Screw, Lap Bipolar Mariland, Lap Toothed Grasper, 10 mm Trocar with Port, 5mm Trocar With Port, Veris Needle 9 inch, Veris Needle 8 inch, Lap Reducer 5 mm, Lap reducer 10 mm, HYS Grasper. |  |  | MAJOR OT
34 | IVF035 | Ultrasound Examination Couch at Gandhi Hospital | 1 | 1 | Nos | 81250 | 81250 |  |  | 202305241093 | 1. Holes of 1 cm dia to allow fluid drainage.2. Lower frame and intermediate frame of steel tubes of rectangular and squareSections, multiple pre-treated and epoxy powder coated.3. Size (approximate): Length: 2050 mm bed surface, 2125 mm with frame.Width: 7504. mm. Bed surface size: 705 mm W x 1950 mm L. 5. Mattress – High-density foam mattress anti-microbial treated with Waterproof flame retardant, antimicrobial, leather like upholstery.6. Backrest operated by gas system from horizontal to seated position _15ºTrendelenburg 7. Height adjustable leg rests 8. Trendelenburg position operated by gas system 9. Pre acceptance demonstration of furniture is must. 10. Motorised with 3 functions- Head rest, TR, High/low 11. Doctor stool 12. Step stool |  |  | ULTRASOUND ROOM
35 | IVF094 | AIR SHOWER Supply and Installation, testing and commissioning of 16 G GI Powder coated suitable single person entry air shower | 1 | 1 | Nos | 355000 | 355000 |  | Ahlada |  | As per Annexure - I
36 | IVF095 | Supply, installation, Testing and Commissioning of Static Pass Boxes of Size 450X450mm with 1.2mm thick SS 304 Mat finish , with interlock doors, UV light On when doors closed condition. As per clean room standard and requirement | 3 | 3 | Nos | 61000 | 183000 |  | Ahlada |  | Supply, installation, Testing and Commissioning of Static Pass Boxes of Size 450X450mm with 1.2mm thick SS 304 Mat finish , with interlock doors, UV light On when doors closed condition. As per clean room standard and requirement.
37 | IVF034 | Mulitfunctional Printer | 1 | 1 | Nos | 43750 | 43750 |  | HP |  | CNB1Q315P2 |  |  | CONSULTATION ROOM
38 | IVF037 | Laptop | 2 | 2 | Nos | 115000 | 230000 |  | HP |  | 5CD324NHNC, 5CD324NGFY |  |  | EMBRYOLOGY LAB, CONSULTATION ROOM
AIR HANDLING UNIT
1 | IVF0151 | Modular Double skin floor mounted EUROVENT Certified Air Handling Unit 3400CFM / 8.5 TR / 150mm SP made out of 45 plus or minus 2mm thick PUF insulated panels Thermal break aluminium profile frame work with 0.8mm thick pre painted / powder coated inner and outer skins comprising of mixing box with fresh air and return air dampers, pre filters with 10 and 5 microns combination filters in single frame, Interlaced/Intertwined DX Cooling Coil suitable for connecting 2 nos of 8.5TR capacity 1W plus 1S out door condensing units, SS304 drain pan with , blower section with Nicotra / Kruger make direct drive/ Plug type fan, vibration isolators, EFF1 motor and blower section with on/off switches to operate from outside, limit switch for the fan section access door with a view port, fine filter section with 3 microns filters inspection doors for all filter and fan sections and damper at the outlet of unit. Ports for differential pressure gauges across all filters and blower to measure differential pressure. AHU Make : VTS / Systemair, Make - Blower - Nicotra / Kruger ; Motor - Crompton Greaves / Siemens ; Filter - Thermodyne / Freudenberg / Ultrafil at Gandhi Hospital | 1 | 1 | Nos | 590000 | 590000 |  | VTS |  |  |  |  | IVF OT & Embriology lab
2 | IVF0166 | SITC of Magnehelic gauges across pre and fine filter at AHUs including mounting arrangement, SS nozzles, food grade PVC tubing, etc. At Gandhi Hospital | 2 | 2 | Nos | 10500 | 21000 |  | Dwyer,Equity make |  |  |  |  | IVF OT & Embriology lab
3 | IVF0167 | SITC of Magnehelic gauges to measure the room differential pressures in operation theatres with SS mounting box, SS nozzles, food grade PVC tubing, etc. At Gandhi Hospital | 3 | 3 | Nos | 12900 | 38700 |  | Dwyer,Equity make |  |  |  |  | IVF OT & Embriology lab
4 | IVF153 | SITC of air cooled condensing units with two scroll compressors, air cooled condenser, power panel, vibration isolating pads, stand for ODU, thermostat, etc. of the following capacities. The nominal capacity of the condensing unit shall be as under. Air cooled condensing unit shall have horizontal discharge. All refrigerant piping accessories sight glass, drier, expansion valve, etc. shall be supplied with the condensing units and suitable to connect above AHUs. Refrigerant shall be R 407C / R 410A. Necessary thermostat, control wiring, etc. shall be included. Derating of the unit shall not be more than 15percentage under 45 deg. C ambient temperature. Condensing unit shall be capable of operating up 50 deg. C ambient temperature. 8.5 TR capacity At Gandhi Hospital | 2 | 2 | Nos | 182000 | 364000 |  | Voltas/Blue Star |  |  |  |  | IVF OT & Embriology lab
5 | IVF154 | SITC of Flange type, Flame proof FLP duct heaters of 5 kW capacity with Mineral Filled AISI SS304/SS321 Sheathed Tubular Heater with GI Helical fins, Tube dia 8.2mm min and Helical Finn dia 25mm min, enclosed in the frame, FLP Terminal box etc., Nickel Chromium Inner Element, Silicon Sealing to Restrict Moisture pickup, Silver brazed mounting plate etc., and heaters shall be controlled with PID thyrestor controller based on the analog output signal of 0-10 V DC of controller . And this should have saftey thermostat and cutouts and air stat etc. all safety controls. The capacities as follows and sizes of the heaters shall be as per the duct sizes and approved drawing. at Gandhi Hospital | 1 | 1 | Nos | 205000 | 205000 |  | Matrusree |  |  |  |  | IVF OT & Embriology lab
6 | IVF152 | Supply and Installation of AL-powdercoated bottom loading type HEPA filter modules with bewel gear arrangement damper with access from clean room suitable to house the following size flange type HEPA filters including GI pressure plate, GI cleats with bolts and pressure drop & Integrity testing ports into room. 600 X 600 X 150MM - HEPA filter - Module height <=350 mm, 600 x 600 x 150 mm - 1000 CFM | 6 | 6 | Nos | 42000 | 252000 |  | AFF/Ulterafil |  |  |  |  | IVF OT & Embriology lab

## Sheet1
6.1.Single Dome, Major
6.2.Intensity control : continuous (1,60,000 Lux)
6.3.Height Adjustment : 600mm
6.4.Action radius : 1850mm
6.5.Possible movements : Radial, Angular and Axial